## Slide 1: Control #94
Passive Noise Suppression Methods for Improvement of Photoacoustic and Photothermal Measurements of Aerosol Light Absorption from Mobile Platforms
W. Patrick Arnott,1 Arthur J. Sedlacek,2 John M. Hubbe3 1Dept of Physics and Atmospheric Science, University of Nevada Reno, Reno NV,
2Brookhaven National Laboratory, Environmental Sciences Department, Upton, NY,
3Pacific Northwest National Laboratory, ARM Aerial Facility, Richland, WA
Part of the Biomass Burning Observations Project (BBOP) for the Department of Energy

## Slide 2: BC and Climate (from Bond et al, 2013 JGR)

## Slide 3: Flight 13 August 2013 Highlights Project Goals
Note the strong difference in the smoke color.
MODIS Image From NASA
How does the smoke optics and chemistry evolve downwind?

## Slide 4: Simulation of Smoke and Air Pollution
Fine Aerosol Optical Thickness is a measure of the smoke in the atmosphere.
It is what satellites ‘see’. It is related to aerosol concentration and health concern.

## Slide 5: Dramatic GOES Satellite Imagery of Fire Dynamics

## Slide 6: Another Example: Reno Smoke 18 Aug 2013

## Slide 7: Biomass Burn Observation Project (BBOP)
L. Kleinman and A. J. Sedlacek (BNL)
Pat Arnott (U. Nevada)
P. Buseck (Arizona State Univ.)
P. Davidovits (Boston College)
E. Lewis (BNL)
W. Lin (BNL)
H. Moosmüller (Desert Research Institute)
T. Onasch (Aerodyne Research)
J. Shilling (PNNL)
J. Wang (BNL)
R. Yokelson (U. Montana)
R. Zaveri (PNNL)
GOES images of Pony Express fire: 08-13-2013

## Slide 8
Biomass Burning
C. Ichoku and R. Kahn
Global Fire Activity MODIS Fire detections for 10 days: Apr 30 – May 9, 2012

## Slide 9
Black carbon exerts positive aerosol forcing (warming) - second only to CO2
Aerosol type can vary
  Flaming phase favors BC production
  Smoldering phase favors organic aerosol
  Fuel source
BB is a significant source of brown carbon (BrC)
  Exhibits pronounced λ dependence in absorption
  Role at CCN (in contrast to BC)
Total climate forcing due to BB is estimated to be:
-0.11 (-0.46 to +0.15) W m-2 (Bond et al. 2013)
  Uncertainly reflects knowledge gaps in BC-cloud interactions & BC interactions with co-emitted organic carbon
Biomass Burn Aerosols
Bond et al. 2013
BC dominate
POA dominate
Notes: Remember to point out that comparisons with satellite observations indicate that bottom-up estimates of aerosol emitted from biomass burning are too low by factors of two to four times which also suggests that emission inventory contributes to this uncertainty.

## Slide 10
Scientific Challenge:
To understand and quantify the role of BB in aerosol forcing (heating/cooling)
  Investigate the impact of particle morphology and chemical composition on BB particle light absorption
  Better understand the microphysical processes that dictate particle morphology (and hence the optical properties) of BC-containing particles
  Examine and quantify how BB aerosol properties evolve in the near field
  What is the minimum knowledge of chemical composition, source, and particle morphology necessary to accurately parameterize the contribution of absorbing aerosols to radiative forcing?
Radiative Forcing Contribution of Biomass Burn (BB) Aerosols
Laboratory scale
Near-field scale
Meso scale
Flame lab/MSU
MODIS
100 km
Notes: In order to bound the radiative forcing contribution of BB aerosols fundamental research into both the properties of BB aerosols as well as the processes that dictate those properties must be undertaken. As presented earlier, the best estimate for total BB radiative forcing is negative (cooling) the uncertainty spans the other side of the fence. Therefore, further research is needed to better bound the sign and magnitude of BB radiative forcing. We will contribute to this effort by focusing in on investigating BB aerosol properties in the near field and how these properties evolve in the near field. By near field we mean within the first couple of hours of aerosol aging. We will be investigating the impact of particle morphology and chemical composition on BB light absorption. As I discussed this morning, our work on the negative lagtimes showed pronounced correlation with biomass burn tracers. The logical question is if BB events favor the creation of near surface rBC inclusions, how does this type of particle morphology impact BB light absorption. Similarly we will want to examine the question of how the chemical composition exerts influence over light absorption in a broader effort to begin addresses the question of relative contribution of particle morphology and chemical composition to aerosol light absorption. Hand and hand with this is understanding the microphysical processes that control particle morphology. BB events are unique in that at the source particle coagulation would be expected to be a major contribution to particle growth. As with our laboratory expts, BB BC-POA coagulation could led to BC inclusions on the surface. But much is going on. What is the role of condensation is such an emission rich environment? Similarly, the role of evaporation. And given this environment, how do the properties of BB aerosols evolve? By answering such questions, we can begin to bound the amount of knowledge necessary to accurately capture reality.

## Slide 11
PI’s – Larry Kleinman (BNL) and Arthur J. Sedlacek (BNL)
+ ~ 30 Scientists from 12 National Labs, Universities, and Companies
120 hours of G-1 Flight Time
July – September: Wildland fires in the Pacific NW
October: Prescribed agriculture burns in the
lower Mississippi River Valley
17 wildfires sampled (ID, OR, & WA)
3 dozen agricultural burns sample (AR)
7 urban plumes (emission contrast)
Biomass Burning Observation Project (BBOP)
Notes: Using the IMPROVE network (Interagency Monitoring of Protected Visual Environments) Park et al. (2007) concluded that in the mid-northern latitudes that BB are a major source of aerosol mass and that year-to-year variability of this mass is dictated by these events.

## Slide 12
G-1 Platform
S.Collier and S. Zhou

## Slide 13
G1 was outfitted with two cameras: forward-looking and Nadar
Scenes from BBOP: Wildfires
Government Flats fire (2013-08-21) was one of the larger fires sampled during BBOP. Over the course of 24-hrs, size increased by > 2500 acres
Org loading: ~ mg/m3
rBC loading: 10s-100s μg/m3
Bscat ~10000 Mm-1 and Babs ~ 1000 Mm-1
Nadar
Forward-looking
100 m
Colockum fire
Government Flats

## Slide 14
Scenes from BBOP: Agricultural Burns
Nadar
Agricultural burns provide distinct advantage of being a single source burn. However, fire lifetime (flame/smoldering phase) much shorter duration than wildfires
MODIS no help in identifying potential sites. Most efficient approach: fly high and look fires
2013-10-20: largest Ag-burn sampled.
Forward-looking
Notes: The instrument suite is uniquely targeted for measurement of BB. To realize this, several new or recently upgraded instruments were deployed.

## Slide 15
This field campaign will leverage the capabilities of several new instruments or instrument combinations that have not been previously used in aircraft.
Microphysical Properties:
  SP-AMS
  FIMS
  Microscopy (TEM)
  SP2
  Dual column CCN
  UHSAS/PCSAP
  Particle counter
Trace gas
  PTRMS
  H2O, CH4, N2O, NO, NO2, NOy, CO, CO2, O3 and SO2
Optical Properties
  3-λ nephelometer (scat; 450, 550, & 700 nm)
  3-λ PSAP (abs; 461, 523, & 648 nm)
  1-λ PAS (abs; 355 nm)
  1-λ PTI (abs; 532 nm)
  1-λ CAPS (ext; 628 nm)
Radiation
  SW, Upwelling hemispheric, spectral
  SW, Upwelling hemispheric, broadband IR. Surface Temperature SW, Down-welling hemispheric, broadband, global and diffuse
  SW, Down-welling hemispheric, broadband, diffuse
BBOP Instrument Suite
Notes: The instrument suite is uniquely targeted for measurement of BB. To realize this, several new or recently upgraded instruments were deployed.

## Slide 16
Several new/upgraded instruments are being deployed during BBOP
New instruments Not Previously Deployed on an Aircraft
Trace gas
SP-AMS
FIMS
TEM Sampler
CAPS
SP2
PAS
PTI

## Slide 17
“28 Mile Marker” Fire sampled at source and 1, 2, 3, 4 hour downwind
Repeatable
Rapid increase within 1st hour in SOA (25%) and scattering (50%)
SOA Increase
BBOP: Near-Field Evolution of Smoke Aerosol Properties
Soot Particle-AMS, Onasch
Scattering increase
0 1 2 3 4 hours
Pasco
Fire
Fire
0
1
3
2
4
hours
PASCO
Terra MODIS image
20 km
Targeted 07-26-2013
Notes: Our goal is to measure the evolution of BB aerosol properties.
Data shown here shows that we are successfully capturing this evolution.
Cite the SOA angle here.

## Slide 18
“Colockum Tarp” Fire sampled at source and 1 - 5 hours downwind
BBOP: Near-Field Evolution of rBC Mixing State
Rapid increase in fraction thickly coated rBC
Targeted on 2013-07-30
Fire
Wind
5 km
Notes: Here we are looking at mixing state and not morphology. Happy to report that negative lagtimes have been observed in BBOP, but not all the time. Need to understand why.

## Slide 19
Quantify the downwind time evolution of microphysical, morphological, chemical, hygroscopic, and optical properties of aerosols generated by biomass burning
Use the time sequences of observations to constrain processes and parameterizations in a Lagrangian model of aerosol evolution
Incorporate time evolution information into a radiative transfer model for determining forcing per unit carbon burned.
BBOP: Relevance in Community
BBOP bridges the gap between laboratory studies (nascent particles) and previous field campaigns of BB (older particles)
Notes: Point out that the community has already expressed interest in BBOP. BNL, in collaboration with Mary Gilles (LBL) recently submitted a proposal to ALS to conduct STXIM analysis on BB particles. This will nicely complement the TEM analysis. I should also point out that we have been quite thoughtful about our field measurements. For example, in several missions, we have conducted back-to-back transects. This means that we have two different grids that can be utilized by two different analysis at two different on the same plume transect and same age. Similarly, U. Arkansas Prof. Gaffney has a Ph.D. student who will be modeling fires and has already expressed interest to use BBOP data.

## Slide 20: Primary Study Region Pacific Northwest

## Slide 21: Pasco WA Airport

## Slide 22: Richland WA AERONET Site

## Slide 23: Fires: Satellite Imagery (blue circle is Reno: Red are fires)

## Slide 24: Some of our Previous Measurements of Biomass Burning Aerosol Optics

## Slide 25: Photoacoustic Measurements from the Twin Otter

## Slide 26: Vertical Profile 27 May 2003: Pancaked Layers of Russian Forest Fire Smoke
Significant layers of light absorbing aerosol near 650 mb and 550 mb, with a clean layer in between at 600 mb.
PHOTOACOUSTIC LIGHT ABS.

## Slide 27: Fire Science Laboratory in Missoula Montana

## Slide 28: Fire Science Laboratory
Instrumentation sampling wood smoke during FLAME Chamber burns
Photoacoustic instruments

## Slide 29: Fuels Tested at the Fire Science Laboratory
29
Rice Straw
Ponderosa Pine
Chamise
Notes: 29

## Slide 30
Key Aerosol Optical Parameters
Fraction of total light extinction (Babs+Bsca) due to
scattering,
Completely scattering ,e.g. sulphate B1
Strongly absorbing diesel soot, freshly emitted BC,  W0.3
Depends on chemical composition and particle size and morphology
Notes: 30

## Slide 31
Key Aerosol Optical Parameters
Angström Exponent for Absorption.
 ≈ 1 for motor vehicle emission generated BC.
Biomass burning aerosols exhibit  as high as 3.5.
Depends on chemical composition and particle size and morphology.
Use measurements at 405 nm and 870 nm to determine 
Notes: 31

## Slide 32: 1,000 Fires, Northern CA, Summer 2008: Peak Smoke in Reno around local noon, July 10th.
Summer wild Fires
-Over 1,000
-Most sparked by lightning
-Dry and windy
-Covering 680 Sq miles of central and northern California
32
Source: NASA California Wild Fires 2008
Notes: 32

## Slide 33: Reno Days in Photographs: Similar Location and Sun Angle
JULY-10, 2008
SEPT-22, 2008

## Slide 34: One Not So Fine Day in Reno

## Slide 35: Chemistry of Primary Smoke Emissions Varies Greatly Too!!!!! (Aerodyne Aerosol Mass Spec measurements.)

## Slide 36: Urban and Biomass Aerosol
They are vastly different!!!!!!!!!

## Slide 37
Chart: 
| 0.866710891413841 | 0.869793966707346
Ponderosa Pine, mixed fuels | 1.7116 | 1.6513
Chamise | 1.1875 | 1.3237 | 1.1974
Rice Straw | 2.8091
Ponderosa Pine Duff | 2.7929
Alaskan Duff | 3.3763
So. Californian Manzanita | 1.1828
Sage & Rabbitbrush | 1.1608
Lodgepole Pine, mixed fuels | 1.8162
Utah Juniper, foilage/twigs | 0.98763
Puerto Rico Fern | 2.6892
Wax Myrtle, branches/foilage | 2.1628
Southern Pine, needles | 1.7205
Comparison With Lab Data
37

## Slide 38: DOE AAF G1 Aircraft

## Slide 39: Aircraft Inlet System
Inlet isokinetic conditions are maintained by adjusting outflow with a mass flow meter
Inlet
Outflow
Expansion slows flow

## Slide 40: DOE G1 Instrument Photos

## Slide 41: Photoacoustic Instrument Installation on Aircraft Rack
Main Part of Instrument:
Acoustic resonator,
Interface board, etc.
Scattering Preamp
Microphone Preamp
Laser control
Computer and DAQ
Aircraft tire innertube
Aircraft tire innertube

## Slide 42: Photoacoustic Instruments For Light Absorption Measurements
Basic principles:
Laser light is power modulated by the chopper.
Light absorbing aerosols convert light to heat - a sound wave is produced.
Microphone signal is a measure of the light absorption.
Light scattering aerosols don't generate heat.
Listen to the sound made by air pollution heated by light.

## Slide 43
Aerosol Absorption, Scattering,
Extinction Coefficients. Single Scatter Albedo.
355 nm BBOP
405 nm,870 nm
532 nm
671 nm, 676 nm
1047 nm
Aerosol Optics Instrument
f0
f0 + df
Notes: 43

## Slide 44: Single Laser Photoacoustic Instrument: Equation to Obtain the Light Absorption Coefficient

## Slide 45: Photoacoustic and Photothermal Rack

## Slide 46: Sound Spectrum During Flight on 10 July 2013

## Slide 47: Sound Spectrum Time Series in Cabin During 10 July 2013 Flight
The sound spectrum changes during flight. Note the strong selected signals.

## Slide 48: 10 July 2013 Test Flight
Flew the Columbia River Gorge to Portland
Some legs were flown through Portland
Intercepted a brief fire plume on the return

## Slide 49: Light Absorption Measurements
Diagnosis: Prior to flight, noise was low. During flight, the noise was very large (red curve) though it was low during taxi. Therefore the inlet system wind noise is the primary source of noise for the system.

## Slide 50: Light Scattering Measurements at 355 nm
Runway prior to flight
Portland, Richland Plumes
EXCELLENT!!

## Slide 51: Post Flight Analysis
Acoustic background in flight was too high!
The main acoustic noise affecting measurements was from the inlet system, analogous to the noise you hear when you stick your head out the window in a car moving at highway speeds.
Remedy:
Try an ‘acoustic resistor’ to reduce low frequency noise.
Construct a ‘muffler’ to try to remove as much noise from the inlet system.
Try not to increase particle loss along the way.

## Slide 52: Acoustic Muffler for Photoacoustic Measurements: First Design

## Slide 53: Acoustic Muffler Prototype: Irrigation tubing and screen, and acoustic foam on the sidewall

## Slide 54: Acoustic Muffler Installed on the Aircraft

## Slide 55: Acoustic Resistor for Low Frequency Damper: Wire Welding Tip with 0.025” Tube Diameter

## Slide 56: 11 July 2013 Flight
Flight purpose: Training for pilots.
Ran the photoacoustic instrument to test its background after I installed the acoustic resistor and the muffler to reduce instrument noise.
Flight was aborted on test landing because of problem with the landing gear.
I was not on the flight.

## Slide 57: Backgrounds and Instrument Pressure
11 July is with acoustic resistor and muffler installed.
11 July shows ambient pressure is about 40 mb lower due to resistance.
Instrument pressure changes as aircraft changes altitude.
Both days, scattering background changes rapidly during the first 15 minutes, then settles.
Absorption background was MUCH more stable on 11 July 2012 after installation of the muffler and acoustic resistor. In this regard the acoustics changes seems to have helped.

## Slide 58: 11 July 2013 Absorption
PROBLEM!!! The noise floor is lower now. However, now have a coherent noise source that raises the photoacoustic signal (red noise is LOWER than the signal now, always!) Suspect that the acoustic resistor is ‘whistling’ to create the coherent background.

## Slide 59: 11 July 2013 Scattering Measurement
Scattering seems to be in excellent shape. Need to study zeros carefully.

## Slide 60: Conclusion from the 10 and 11 July Flights
Scattering measurements are working well.
The laser needs 30 minutes warm up in the beginning to stabilize its pointing stability.
On ground absorption noise background is about 10 Mm-1. It will always go up in flight.
The installed acoustic resistor and muffler reduced the absorption noise measurement by more than a factor of 2; however, I think the acoustic resistor will need to be removed because it seems to be whistling, seems to be creating a coherent noise on the acoustic measurement.
Prior to the next flight, the acoustic resistor should be removed.
Also may want to improve the muffler performance.
May install a noise canceling system for the 1510 Hz frequency used by the PAS.

## Slide 61: 15 July 2013 flight
First research flight to study the Portland outflow plume.
Results:
Relatively low aerosol scattering at all times.
Photoacoustic instrument preamp was not turned on. Absorption not measured.

## Slide 62: 15 July 2013 Scattering Background

## Slide 63: 15 July 2013 Scattering at 355 nm
FLIGHT

## Slide 64: Improved Muffler: Add a bandstop filter

## Slide 65: Muffler Construction in Progress
Screen Cylinder (used also as an irrigation system filter)
Aluminum disk
Radial wave oscillator section
Top and bottom are symmetrical. Walls coated with conductive paint.
Absorptive volume filled with open cell foam to absorb sound
Screen cylinder: 1.67” diameter
Radial oscillator: 4.5” diameter
Nominal 4” sewer system T

## Slide 66: Testing the Muffler
Speaker
Microphone 1
Microphone 2
Sweep the speaker through the frequency range of interest.
Record the ratio of microphone 2 and microphone 1 signals.

## Slide 67: Sound Attenuate of -84 dB at 1500 Hz!

## Slide 68: Reality Check: Sound Transmission by An Equivalent Length of Tubing

## Slide 69: Results of Reality Check: Sound Transmission by An Equivalent Length of Tubing: Oscillates Around 0 dB

## Slide 70: Muffler Installed on Aircraft (installation and photo by John Hubbe)

## Slide 71: 17 July 2013 Research Flight to Biomass Burning Smoke in Idaho
Words of Dr. Jennifer Comstock, PNNL:
“A successful research flight was completed yesterday into several fire plumes over the Salmon River Mountains in Idaho.”
Rough Creek and Ridge fires of Idaho were studied.

## Slide 72: Fire Descriptions

## Slide 73: Flight Track
Rough Creek Fire
Pine Creek Fire
To Richland WA

## Slide 74: TSI Neph and PSAP Data In Flight

## Slide 75: Acoustic Resonator Properties
Bad calibration caused problems.

## Slide 76: 355 nm Absorption: Still Too Noisy!!
Bad Acoustic Calibration
Background Noise TOO HIGH!
Bad Zero Propagates ???

## Slide 77: 355 nm Scattering: Excellent 

## Slide 78: Laser Power and Instrument Temperature 

## Slide 79: 19 July 2013 Flight
PSC to Ridge Fire at 11500 ft.
Flew down to Summit fire, not as pronounced plume.
Turned back toward Ridge Fire and flew downwind of plume perpendicular to plume flow.
First leg @13000 ft.
Next leg at 10,000 ft 5 miles downwind from first leg.
Next leg at 11,000 ft centered on Ridge plume, reversed and did same line at 12,000 ft. Last leg at 13,000 ft.
Pine and Ridge fires were raging. Went 22 miles downwind from fire. Circled around and did the pattern again only in the plume of the ridge fire. Last pass was through the column of the main plume for Ridge fire (thick dark black smoke).

## Slide 80: Ridge and Summit Fires (from http://www.idahostatesman.com/2013/07/18/2659851/ridge-fire-grows-more-than-600.html)

## Slide 81: Flight Track: Pasco to Central Idaho

## Slide 82: Light Absorption (old muffler) Very noisy data in flight

## Slide 83: Light Scattering Data Good

## Slide 84: State Parameters: Temperature Went high for awhile in flight

## Slide 85: 23 July 2013 Flight
AAF conducted two crew-training flights today, both doing sampling at the Sunnyside Turnoff fire in Central Oregon, with partial payload.
TURNED OUT TO BE A GOOD DAY FOR PHOTOACOUSTICS 
Fire View from Tarmac

## Slide 86: Light Scattering And Absorption at 355 nm
FIRE PLUME PENETRATIONS
USEFUL BABS DATA 

## Slide 87: 25 July 2013 Flight
Sunnyside Turnoff Fire and Mt Batchelor Observatory Overflight

## Slide 88: 25 July 2013 Flight
Sunnyside Turnoff Fire and Mt Batchelor Observatory Overflight
Also sampled this fire.

## Slide 89: 25 July 2013 Flight
Sunnyside Turnoff Fire and Mt Batchelor Observatory Overflight
Pretty thick smoke plumes with dark aerosol.

## Slide 90: 25 July 2013: First flight with the Muffler Model 2 (installation and photo by John Hubbe)

## Slide 91: Muffler 2 doesn’t help much!
Noise Equivalent Absorption is MUCH lower with muffler 1 than muffler 2.

## Slide 92: 25 July 2013 Scattering
Some large scattering signals were observed during various plume passages.

## Slide 93: 25 July 2013 Resonator Data
All looks well.

## Slide 94: BUT !!!!
Turns out, on the 25th the muffler and inlet system was not set up properly. Therefore we will have to wait until later to find out what is going on.

## Slide 95: 26 July 2013: Fire in South Central WA

## Slide 96: 26 July 2013: Fire in South Central WA
Looks like a fairly low single scattering albedo: actively burning fire, maybe brush.
EXCELLENT coverage of the fire with many plume intercepts.

## Slide 97: 355 nm Absorption with Muffler 2

## Slide 98: Muffler Comparison
Muffler 1
No Muffler
Muffler 2
Key Messages:
Mufflers help reduce acoustic noise (see red curves on each.)
Mufflers 1 and 2 are superficially similar in their effectiveness

## Slide 99: Muffler Comparison: More Details
Muffler 1
Muffler 2
Key Messages:
Mufflers help reduce acoustic noise (see red curves on each.)
Mufflers 1 and 2 are superficially similar in their effectiveness
Muffler 1 resulted in severe noise in the past, at times, that made for a bad PAS acoustic calibration.
Next slide compares ambient pressure (higher likely more turbulent).

## Slide 100: Muffler Comparison: Effect of Aircraft Altitude (measured by instrument ambient pressure)
Key Message: Aircraft altitude is related to location in the boundary layer, and effective turbulence. It was similar for both days; therefore we can more or less directly compare noise performance.

## Slide 101: Direct Comparison of Noise for Mufflers 1 and 2
Noise performance was similar for the mufflers. Muffler 2 (26 July 2013 data) had more constant values (could be still differences due to turbulence, etc).

## Slide 102: Histograms of Noise From Both Mufflers
Message:
Average Noise (> 50 Mm-1) Muffler 1 = 174 Mm-1
Average Noise (> 50 Mm-1) Muffler 2 = 187 Mm-1
Still prefer muffler 2 for now as it was associated with no bad zeros.

## Slide 103: 355 nm Scattering 26 July 2013
Very strong plumes were observed. Scattering background was excellent.

## Slide 104: 30 July 2013: Colockum Fire Near Wenatchee WA

## Slide 105: Bsca Components
Scattering measurements are working well. The laser power is dropping. It is likely that the instrument windows may need changed or cleaned soon.

## Slide 106: Babs Components
Absorption Signals are Still Noisy!

## Slide 107: Filtered Data 355 nm: Bsca > 100 Mm-1 Babs = Babs + BabsBackgnd – 35 Mm-1
Babs comes out of the noise here

## Slide 108: 6 August 2013 Flight
Coordinated flight with NASA.
Installed both mufflers on the sample line for this flight to test their effectiveness.
Sampled fires in SW Oregon.

## Slide 109: Both Mufflers 1 and 2 have been Installed
Muffler 1
Muffler 2

## Slide 110: Muffler Comparison: More Details
Muffler 1
Mufflers 1 & 2
Key Messages:
Mufflers help reduce acoustic noise (see red curves on each.)
Mufflers 1 and 2 together are more effective than just muffler 1
Use of both mufflers resulted in continuous measurements during the entire flight
Will install muffler 3 in place of muffler 2, and will try the acoustic resistor on the inlet along with sound insulation on the PAS laser cavity

## Slide 111: Muffler and Acoustic Resistor Comparison
Muffler 1 + Acoustic Resistor
Mufflers 1 & 2
Key Messages:
Looking back to the 11th of July, I misinterpreted the effectiveness of the acoustic resistor.
The noise spectrum with the acoustic resistor and muffler 1 is lowest (see red curves).
Will install muffler 3 in place of muffler 2, and will definitely put the acoustic resistor on the inlet along with sound insulation on the PAS laser cavity. The acoustic resistor should be in front of the first muffler, closest to the inlet.

## Slide 112: Comparison of Babs and Bsca
BabsNew = BabsOld + background – 65 Mm-1 processing of absorption.
100 points averaged.
The main event is resolvable.

## Slide 113: Flight on 9 August 2013
Returned to the Colockum Tarps Fire

## Slide 114: Absorption and Scattering Components
Some very clear spikes were observed.
Absorption needs some smoothing and zero correction.
Babs = (Babs + Background – 60 Mm-1 ) * Q / 65

## Slide 115: Resonator Qualities

## Slide 116: Refined Time Series

## Slide 117: Single Scattering Albedo at 355 nm

## Slide 118: Comparison of Bsca and SSA
Peaks of SSA (whitest aerosol) correspond with peaks of scattering coefficients.

## Slide 119: Transmission Spectra of 3 Mufflers (obtained 12 August 2013 while in Pasco)

## Slide 120: Muffler 3 and Acoustic Resistor Installed on 12 August 2013
PAS
PTI
Acoustic
Resistor
(0.025” diameter,
1” long copper capillary)

## Slide 121: Moved the Acoustic Resistor To My System Only 13 August 2013
Acoustic resistor was found to cause the PTI to have excess noise when the PAS was in the zeroing phase. Now, with the acoustic resistor on my system only, the following pressures were observed. Pambient = 1004 mb. Pinstrument = 967 mb, = 950 mb when zeroing.

## Slide 122: Muffler 3 Construction
4.5” Diameter Sewer Pipe Cleanout
Stainless steel screen is rolled up. The foam is rolled around the screen.
A feed through fitting is used. The ‘dumbell’ on the right contacts the screen.

## Slide 123: Sound Attenuation by Acoustic Resistor

## Slide 124: Acoustic Resistor for Low Frequency Damper: Wire Welding Tip with 0.025” Tube Diameter

## Slide 125: Before Takeoff Procedures
The scroll pump (for the PAS and PTI among others) is turned off prior to take off.
It is turned back on when in the air.
The instrument pressure will therefore drop when the pump goes back on.

## Slide 126: Al Jazeera Interviewing Jennifer Comstock of PNNL Concerning BBOP

## Slide 127: Flight 13 August 2013 Highlights Project Goals
Note the strong difference in the smoke color.
MOSIS Image From NASA
How does the smoke optics and chemistry evolve downwind?

## Slide 128: Dramatic GOES Satellite Imagery of Fire Dynamics

## Slide 129: Babs Components: Huge Signals!
The noise floor is about 100 Mm-1, much improvement with the muffler and acoustic resistor.

## Slide 130: Bsca Components: Excellent!
SSA is about 0.8 during the peak.

## Slide 131: Refined Analysis: Babs, Bsca, and Single Scatter Albedo at 355 nm

## Slide 132: 2 Minute Average Time Series

## Slide 133: Scatter Plot and Histogram Of SSA

## Slide 134: 14 August 2013 Flight
Repeat of yesterday’s flight plan to the massive fire in Southern Idaho.
http://www.independentmail.com/news/2013/aug/13/hundreds-evacuated-amid-idaho-wildfires-some-stay/
Very dark smoke, flaming bush

## Slide 135: More on the Fires In Idaho
http://www.idahostatesman.com/2013/08/12/2702396/elk-complex-fire-burns-close-to.html
Note the darker and lighter smoke

## Slide 136: Absorption and Scattering Raw Data
Very nice looking data. Some plumes were intercepted, especially towards the end.
It was difficult to find the plume core due to intense smoke in the area (couldn’t see the smoke for the smoke). Very useful flight.
The photoacoustic instrument laser can use a good hour or two to warm up in the preflight time. The laser power increases to a maximum during this time.

## Slide 137: Resonator Performance
All of the acoustic calibrations are very good. The Q and calibration pressure are excellent.

## Slide 138: Refined Babs and Bsca Data

## Slide 139: Single Scatter Albedo
Single scatter albedo is around 0.9 for this flight, similar to yesterday.

## Slide 140: Comparison TSI Neph Extrapolated Scattering and PAS Scattering at 355 nm
EXCELLENT! Some residual difference may be due to averaging difference.

## Slide 141: Time Series of PSAP and TSI Neph Absorption and Scattering: 2 min ave

## Slide 142: Histograms of PSAP and TSI Neph SSA

## Slide 143: Comparison SSA PSAP and PAS
355 nm Results

## Slide 144: 16 August 2013 Flight
Overflight Mt Batchelor Observatory. Likely path shown in red. Forecast back trajectories show winds from the southwest, carrying smoke from California, over Mt Batchelor.
Radar (green) shows an active front NW of the flight track.

## Slide 145: 16 August 2013 Flight
Overflight Mt Batchelor Observatory. Likely path shown in red. Forecast back trajectories show winds from the southwest, carrying smoke from California, over Mt Batchelor.
Radar (green) shows an active front NW of the flight track.

## Slide 146: Scattering Components
Modest aged smoke plumes were observed.

## Slide 147: Absorption Components
Generally low levels; will need further averaging to study the structure.

## Slide 148: Resonator Components

## Slide 149: 21 August 2013: Government Flats Complex
Discussion date: 29 August 2013.
Morning flight was lackluster.
Afternoon, the smoke punched through the boundary layer, and made a spectacular show!

## Slide 150: 21 August 2013 Government Flats Fire Oregon

## Slide 151: Photos From Art Sedlacek During Flight

## Slide 152: Raw Scattering and Absorption: Huge Numbers!!

## Slide 153: Resonator Components: Good except before 1st flight

## Slide 154: Scattering And Absorption 21 August 2013