## Slide 1: Fast Response (< 1 sec) Thermistor Temperature Sensor
By Pat Arnott
Atmospheric Science Program
University of Nevada Reno
3D printing by Jayne Boehmler, Atmospheric Science Grad Student
Circuit board machining Wade Cline, Physics Dept Machinist
All design files and modeling files included in project outcome.

## Slide 2: Motivation for the Temperature Sensor: The Perlan Project – Go for the World’s Record Altitude Glider Flight.
From http://pubs.rsc.org/en/content/chapterhtml/2015/bk9781782620761-00001?isbn=978-1-78262-076-1
The Perlan 2 glider aircraft will ride atmospheric gravity waves, made stronger by interaction with the Antarctic vortex, to a record height for a glider ≈ 30 km altitude. Temperature measurements allow for mapping of the wave structure to characterize it, and help navigate.
http://www.perlanproject.org/aircraft

## Slide 3: Based on the Amphenol Thermistor BR11KA152M (available through Mouser)

## Slide 4: Based on the Amphenol Thermistor BR11KA152M (available through Mouser)
Guides Circuit for Measuring Thermistor Resistance

## Slide 5: Nominal Temperature Table and Stein Hart Equation: Note R(T=25 C) = 1.5 kOhm
where for that table, T is in Celcius, R is in Ohms, and
A= 1.051740553e-3,
B= 3.042515034e-4,
C= 1.974195817e-7.
The website for the Stein Hart calculator is http://www.thinksrs.com/downloads/programs/Therm%20Calc/NTCCalibrator/NTCcalculator.htm
Stein Hart Coefficients were calculated from the website below, using the table values for resistance at -80 C, -40 C, and 25 C.

## Slide 6: Design Using a Voltage Divider Circuit, Thermistor Resistance, and Fixed Resistance = 1 Mohm.
Fixed resistance = 1 MegOhm chosen to keep thermistor power dissipation below 0.007 mW.

## Slide 7: Design Using a Voltage Divider Circuit, Thermistor Resistance, and Fixed Resistance = 1 Mohm.
Maximum thermistor voltage is about 1 volt. The AD623 instrument amplifier provides gain of 4.834.
The ADS1115 analog to digital convertor was chosen because it provides programmable gain of
1, 2, 4, 8, and 16, so that it can help provide good measurements at warmer temperatures where the voltage is about 8 mv.

## Slide 8: Sensor Head Schematic: Note the thermistor is sandwiched between 500 kOhm resistors

## Slide 9: Sensor Head Board Layout Dots are 0.1” apart
Top side
(ground plane)
Bottom side
(5v plane)
Yellow is from top side
stencil

## Slide 10: Actual Board

## Slide 11: Receiving Board Layout Sensors: Digital and analog pressure, GPS, Digital temperature and RH, Thermistor, 5 analog inputs for Teensy Teensy 3.6 microcontroller with real time clock and microSD card data storage. Nokia 5110 display screen Power from 12 volts in, converted to 5 volts. Battery backup for the real time clock

## Slide 12: Transceiver on Receiving Board Purpose: Convert TTL serial output to RS232 standard for old style serial port on computer

## Slide 13: Control Board Layout: Front Side

## Slide 14: Control Board Layout: Back Side (note capacitors C1 through C5)
Yellow stencil is from the front side

## Slide 15: 3D print box for circuit board

## Slide 16: 3D print box top: Slot is for wires from thermistor and digital T and RH sensors

## Slide 17: Board and Sensor

## Slide 18: Theoretical Time Response:
Based on the paper: N. K. Wagner, “Theoretical Time Constant and Radiation Error of a Rocketsonde Thermistor” Journal of Meteorology, vol. 18, 1961, pgs. 606-614.
Figure 1 of the paper is shown here. We will use the theory to an altitude of 25-30 km.
The thermistor in this paper is ventilated as a sensor attached to a falling parachute released when the rocket reaches peak altitude. Note that parachute fall speed increases with height, while time constant increases as well, making the measurement all the more difficult.

## Slide 19: Theoretical Time Response:
Based on the paper: N. K. Wagner, “Theoretical Time Constant and Radiation Error of a Rocketsonde Thermistor” Journal of Meteorology, vol. 18, 1961, pgs. 606-614.
Figure 2 of the paper is shown here. We will use the theory to an altitude of 25-30 km.
The thermistor in this paper is ventilated as a sensor attached to a falling parachute released when the rocket reaches peak altitude. Note that parachute fall speed increases with height, while time constant increases as well, making the measurement all the more difficult.
The Perlan aircraft moves much faster (horizontally) than the parachute falls, so time response calculations are upper estimates.

## Slide 20: Theory
Implemented in Fortran, with subroutines I wrote years ago for the thermal physical properties of air.
Atmospheric temperature and pressure vs height taken from this simple description:
https://www.grc.nasa.gov/www/k-12/airplane/atmosmet.html

## Slide 21: Model Atmosphere

## Slide 22: Our Thermistor Calculated Time Response as if it were falling with a parachute
Unrealistic above
25 km.
Convective heat transfer coefficient needs change. (See fig. 1 of paper for more correct shape).
Notes: Reasonably matches product claim of 0.8 seconds in still air at the ground. Since aircraft speed is probably an order of magnitude larger than parachute (at least), theory is providing an upper limit for the time response.

## Slide 23: Our Thermistor Calculated Time Response as if it moving at 10x fall speed of parachute.
Notes: Time response calculation for speed a factor of 10 faster than a falling parachute.

## Slide 24: Overnight Testing Set Up

## Slide 25: Testing in Ambient Conditions: Overnight

## Slide 26: On Road Set Up

## Slide 27: Testing Driving in Car along I80 in Reno using GPS on control board for speed sensor

## Slide 28: Freezer Test: Suddenly put Thermistor in Upright Freezer.