## Slide 1: Soot Particle Aerosol Mass Spectrometer: Development, Validation, and Initial Application
T. B. Onasch,A. Trimborn,E. C. Fortner,J. T. Jayne,G. L. Kok,L. R. Williams,P. Davidovits, and D. R. Worsnop
By Gustavo M. Riggio
05/12/2014

## Slide 2: Introduction
+
Single Particle Soot Photometer (SP2)
Aerosol Mass Spectrometer (AMS)
Developed to measure the chemical and physical properties of particles containing
black carbon (rBC)

## Slide 3: Introduction
Portable
Real time
Highly sensitive
Expensive

## Slide 4: Refractory Black Carbon (rBC)
Black Carbon (BC)
  Generated by incomplete combustion of fossil fuels, biomass, and biofuels.
  Affect air quality, human health, and direct and indirect radiative forcing.
  Detailed effects of BC highly uncertain.

## Slide 5: Instrument Utility/Development
Single Particle Soot Photometer
  Quantify rBC by detecting incandescent signals.
    Non-incandescing materials will scatter light (i.e. organic coatings)
Notes: Provides rBC mass loadings, particle number concentration , and size distribution
Explain laser beam, incandescence, scattering. (overview from Ben’s presentation)
CHEMICAL INFO MISSING

## Slide 6: Instrument Utility/Development
Aerosol Mass Spectrometer
  Measures composition of nonrefractory aerosol particle ensembles.
Animation of the Aerodyne AMS. Credit: Matt Thyson (Lexington, Massachusetts)
TOF Mass Spectrometer
Notes: Explain entrance region and TOF MS
Lack of sensitivity to refractory particles.
Particles vaporized at 600 C using resistively heated tungsten vaporizer.
Vapor is ionized with an electron beam via electron impact ionization and ions detected in a high resolution mass spectrometer.
Provides average mass spectrometric measurements on a particle by particle basis.

## Slide 7: Instrument Design SP-AMS
Laser ON/OFF
  SP-AMS mode
Chopper OPEN/CLOSED
  MS mode

## Slide 8: Instrument Capabilities
Quantitative detection of black carbon
Information on coatings on black carbon cores
Real time analysis

## Slide 9: Particles Across Laser Beam
Coating evaporates first.
  Low temp. (<600 oC)
Core evaporates last.
  High temp. (> 1000 oC)

## Slide 10: Laser Vaporizer
Ionization efficiency depends on laser alignment
(CCD camera), and power.
Intensity must be sufficient to vaporize particles.
Dispersion of particles may
cause particles to miss the laser.

## Slide 11: Vaporization Overview
Non refractory material vaporizes first.
rBC heats to thousands of degrees.
  Gives rise to visible incandescent signal
Simultaneously, rBC vaporize into carbon clusters.
  Ionized and detected by mass spectrometry.
    AMS not able to vaporize rBC (Filament temp. = 600 oC)
What happens if we turn the laser on and off while the tungsten vaporizer is on? What do we measure?

## Slide 12: SP-AMS Parameters

## Slide 13: Efficiency
Collection efficiency depends on:
  Fraction of particles diverted from laser beam (ES).

## Slide 14: Efficiency
Collection efficiency depends on:
  Fraction of particles lost during transit through inlet and aerodynamic lens (EL).
  Fraction of particles lost due to bounce effects (EB).
CE = EL x EB x ES
AMS Collection Efficiency Issues. http://cires.colorado.edu/jimenez-group/UsrMtgs/UsersMtg9/08_Onash_CE.pdf

## Slide 15: Calibration
Dependent on the measurement of 2 out of 3 variables.
  Relative ionization efficiency
  Mass specific ionization efficiency of a species
  Mass ionization efficiency of nitrate ions

## Slide 16: Calibration…
Ionization Efficiency:
  Ions detected per particulate mass sampled
Relative Ionization Efficiency:
  Ratio of the mass specific ionization efficiencies
10-12 = units conversion
Na = Avogadro’s number

## Slide 17: rBC Calibration
Calibration appears to be dependent on particle type.
  Used Couette Centrifugal Particle Mass Analyzer
    Shape independent measure of particle mass.
Incomplete overlap between particle and laser beam.

## Slide 18: Sensitivity Curve for SP-AMS
Relative rBC ion signal as function of vaporizing laser power.
  rBC reaches a plateau at higher laser power.
  Detection limit not limited by laser power.
Important to operate with sufficient light intensity.

## Slide 19: Sensitivity
See figure S3

## Slide 20: Measure Particulate Species for 3 vaporizer combinations

## Slide 21: Chemical and Physical Information

## Slide 22: Instrument Characterization
Peaks in black are carbon ions.
  Not observed using standard AMS
Provide “finger print” for different combustion sources.
Mass spectrum of denuded ethylene flame soot.
Notes: Change in carbon clusters: below C30 is m/z 12, above C30 is m/z 24 (2 carbon atoms),
Represent changes in structures (i.e linear chains, monocyclic rings and polyciclic, to fullerenes.

## Slide 23: Laser ON/OFF Mass Spectra
Lab generated soot particles
  Laser ON vs OFF
CO2 = largest difference
Same signals may be
present with laser ON and
OFF.
Notes: The CO2+ signal that disappears in the laser-off experiment is not observed in 2 other different types of rBC particles (i.e., the denuded soot shown in Figure 3 or in the mass spectrum of glassy carbon spheres, not shown). Therefore, we conclude that this signal observed in the Figure 4 mass spectrum originates from the particle composition rather than from products of laser induced oxidation of carbon.

## Slide 24: Laser ON/OFF Differences
Sum of the ion signals
  Laser ON vs. OFF
Laser ON – all signals
present
Laser OFF – only organic signals
  Decrease of 20%
CO2 originates from particle composition.

## Slide 25: Coating Effects and CO2
Measures of ion signal
distribution as function of
particle size.
rBC integrated signal
remains the same.
Organic signal increases.
Uneven coating.
Notes: 225 mm size selected – 50mm glassy carbon spheres
Coated via gas to particle deposition of DOS (dyoctil sebacate)

## Slide 26: Ambient Measurements
Spectra dominated
by nonrefractory BC
and inorganics.
Higher C1 – C5 for
ambient than lab.
samples.

## Slide 27: MAAP vs SP-AMS
Good agreement
Organic vs BC
dominated plumes
differentiated
Similar to diesel
exhaust and lubrication
oil spectra.

## Slide 28: Plume Types
Diameter rBC ∼ 120 nm
  Similar in size to diesel exhaust particulate emissions (fresh)
Diameter organics ~ 170 nm
  Consistent with coating effects
Sulfates indicator of the accumulation mode
  Particles least affected by atmosphere (persistent)
rBC from local sources

## Slide 29: Conclusion
Portable, high resolution, real time
Two configurations
  Laser vaporizer (SP-AMS)
  Tungsten vaporizer (AMS)
Provides BC measurements (chemistry, size distribution, and mass loading)
Coating measurements possible