## Slide 1: History of the DMT Photoacoustic Instruments
By W. Patrick Arnott, University of Nevada Reno
Collaborators: Hans Moosmüller, Fred Rogers, John Walker, Barbara Zielinska, John Sagebiel, Eric Fujita, Dave Campbell, Rick Purcell, Dan Wermers, Sebastian Upallalli (DRI), Rich Raspet, Willie Slaton (Univ. Miss.), James Mehl, Rob Abbott, Mike Ossosfsky, Khadeejeh Hamasha, Ali Rahmah, Kristin Lewis, Lupita Paredes, Adel Sarofim, Kerry Kelly, Dave Wagner, Joanne Lighty (Univ. Utah CFE), Otmar Schmid, Meinrat Andreae (Max Planck Institute for Chemistry), John Ogren and Pat Sheridan (NOAA CMDL), John Seinfeld, Rick Flagan, Haf Jonsson (CAL TECH), Anthony Strawa, Beat Schmid, Rich Ferrare (NASA) Manvendra Dubey, Rahul Zaveri (LANL and PNNL) Claudio Mazzoleni (MTU) and Rajan Chakrabarty (WUSL) Ian Arnold (UNR) Darrel Baumgardner, Greg Kok, John Walker, and Gavin McMeeking (DMT) and others … It’s been a wonderful adventure … hopefully continues …
Funding over the Years: EPA, NPS, ONR, NSF, DOE, DOD-SERDP, DRI, UNR

## Slide 2: Introduction
We are not the first to use photoacoustic methods for aerosol light absorption measurement;
(Petzold, Niessner, Adams, Pierson, Japar, Terhune, Anderson, Foot, Truex, Pinnick)
going back to Alexander Graham Bell.
We were at a convergence in history of …
A. Strong need for a practical instrument for aerosol light absorption measurement that improves upon filter-based methods.
B. Technological breakthroughs in lasers to get high power in small, efficient packages.
C. A great opportunity for people interested in aerosol instrumentation development.

## Slide 3: Photoacoustics and Alexander Graham Bell Observations and usefulness of sound from light, Circa 1880

## Slide 4: Arnott’s Postdoc National Center for Physical Acoustics, University of Mississippi: 1988-2001: Study of ThermoAcoustic Heat Engines
Arnott, W.P. H.E. Bass and R. Raspet, 1991: General formulation of thermoacoustics for stacks having arbitrarily-shaped pore cross-sections. J. Acoust. Soc. Am., 90 , 3228-3237. (A key paper for me, and one that was particularly useful to others). Some of the ideas coming out of this time have been realized in a well-used software package for design of acoustic refrigerators and prime movers (sounds sources).

## Slide 5: 1992: Bill Pierson, Director of the Energy and Environmental Engineering Center at the Desert Research Institute
Former Ford Motor Company researcher: Had worked with trailer size photoacoustic instruments.
Strongly advocated for development of practical photoacoustic instruments.
Mentored our EPA Exploratory Proposal submitted (and funded) by Arnott (PI), Moosmüller, and Rogers, 1993-1997.
Arnott
Pierson

## Slide 6: Assembled a team at DRI to pursue Aerosol Light Absorption, Scattering, and Extinction Instrumentation Development
Pat Arnott,
Acoustics and instrumentation
Fred Rogers
Aerosol expert
Hans Moosmüller
Optics and instrumentation
John Walker
Machinist and
instrumentation
Also …
Rick Purcell, DRI Mechanical Engineer
Dan Wermers, DRI Electrical Engineer

## Slide 7: Circa 1992-1993: Learned how to amplify thermally generated sounds with thermoacoustic heat engines

## Slide 8: Obtained enormous signal to noise ratios by thermoacoustic control (very high quality factor for the resonator) Great for pulsed lasers. Needed also to develop a simple system.

## Slide 9: Why photoacoustic instruments? Filter based measurement systems are simple, but much work is needed to interpret results. Realization by the community that alternatives to filter methods were needed.
Aerosol are deposited on the light-diffusing filter, multiple scattering substrate.
Light absorbing aerosol reduce the light power at the photodetector.
Light scattering aerosol don't reduce power in principle.
Calibration and effects of aerosol loading (blocking)??
We have noted large RH induced artifacts (PSAP).
Filter-based systems have low dynamic range because the filters fill up quickly at high concentration.

## Slide 10: PSAP: Particle Soot Absorption Photometer A filter-based measure of light absorption.
Reference filter on the right. Aerosol-loaded filter on the left. Compare light transmission through these filters as a measure of light absorption.

## Slide 11: Aethalometer Filter Details: Side view.

## Slide 12: Photoacoustic Instruments For Light Absorption Measurements
Basic principles:
Laser light is power modulated by the chopper.
Light absorbing aerosols convert light to heat - a sound wave is produced.
Microphone signal is a measure of the light absorption.
Light scattering aerosols don't generate heat.

## Slide 13: Photoacoustic Instrument Details: Equation to Obtain Light Absorption Coefficient.

## Slide 14: Strongly Absorbing Aerosol Response: Note the Ski Slope Response of the Filter Sampler as the Filter Darkens From the Reno Aerosol Optics Study

## Slide 15: Our First Fieldable Photoacoustic Instrument Details: Combined elements of the Sigrist ‘Banana’ Resonator with use of Acoustic Notch Filters to allow windows to largely decouple from the resonator
1” very sensitive microphone
System calibration

## Slide 16: Our First Field Deployment of the Photoacoustic Instrument: North Front Range Air Quality Study (NFRAQS) Winter 1997

## Slide 17: Photoacoustic Instrument at NFRAQS (1997)

## Slide 18: Photoacoustic Data Acquisition (National Instruments Labview).

## Slide 19: Some Results from NFRAQS: 532 nm Laser
Measurements agree at times… time offset, 10 m2/g?
Comparison allows evaluation (calibration?) of filter method for aerosol light absorption measurement.

## Slide 20: 1998 – 2002 NSF Major Research Instrumentation Project Moosmüller PI, Arnott CO PI 1. Further Development of the Photoacoustic Instrument 2. Development of a Low Truncation Angle Nephelometer 3. Development of a Cavity Ring Down Instrument for Aerosol Extinction Measurement

## Slide 21: DOE CARAT PROJECT Simple Particulate Emissions Measuring System
CARAT Cooperative Automotive Research for Advance Technology. (Highly efficient, low-polluting vehicles of the future - from diesel to fuel cell technologies.)
Purpose: Evaluate methods for fast time response measurements (< 5 sec) of aerosol mass from diesel engines.
Background: Vehicles are evaluated using dynamometer test facilities. The EPA standard protocol for aerosol mass measurement use filter samples acquired over many minutes. Instruments and methods for determining aerosol mass at high time resolution should provide automotive, engine, and exhaust system manufactures with valuable insight.
UC Riverside College of Engineering Center for Environmental Research and Technology (CE-CERT) provided the dynamometer test facility.
DRI provided the instruments for measuring high time resolution mass measurements.
Notes: Diesel Engines
Simple Particulate Emissions Measuring System
Current methods for measuring particulate matter emissions in diesel and other internal combustion engines require expensive, complex equipment and are difficult to use. Funded organizations are working to develop a simple, accurate, and user-friendly sensor to directly or indirectly measure particulate matter emissions from internal combustion engines. The goal is to develop a measuring system that is just as accurate but significantly simpler, less expensive, and easier to use than current dilution-tunnel, filtration-based systems. The improved system should be able to measure the most common types of particulates (carbon particles, liquid droplets, sulfate particles) as small as 1 micrometer. A system capable of making continuous (transient) measurements (with updates every 5 seconds, minimum) would be particularly attractive to industry.
CARAT is sponsored by the U.S. Department of Energy through its Office of Advanced Automotive Technologies (OAAT). Unlike other federally funded programs, the three phases of the CARAT Program have been designed to foster and support the development of a technology from proving its feasibility, through prototype development, and on to partnering with a manufacturer for validation of the technology and demonstration of production feasilibity. The program has been established to research, develop, and validate advanced automotive technologies that will enable the production of highly fuel-efficient, low-emission, and fuel-flexible cars and trucks.
For more information on the role of the CARAT Program within OAAT's R&D Plan, click here.

## Slide 22: Dynamometer Test Facility at UC CE-CERT

## Slide 23: Photoacoustic Instrument, 1999, resonator.
Stainless steel resonator, small acoustic enclosure, cleaner design.

## Slide 24: Photoacoustic Instrument, 1999, Laser.
Crystalaser (company in Reno) 532 nm Nd-YAG laser diode-pumped solid state laser. Chopper wheel provides laser power modulation.

## Slide 25: DOE-CARAT Photoacoustic Results (Dodge) Shows the great utility of photoacoustic instruments for black carbon measurements from vehicles.

## Slide 26: PA serial # 1... Built for Adel Sarofim, University of Utah, for Black Carbon Mass Concentration Measurements by Vehicles Designed by Pat Arnott, Machined by John Walker After a Demonstration Project at Hill Air Force Base

## Slide 27: Patent for the PAS 1 Instrument: Single Wavelength Photoacoustic Instrument

## Slide 28: Instrument Description in the Patent

## Slide 29: Site Details: Instrument built for the Max Planck Institute for Chemistry

## Slide 30: SMOCC
Smoke Aerosols, Clouds, Rainfall and Climate:
Aerosols from Biomass Burning Perturb Regional
and Global Climate
S M O C C
http://dionysos.mpch-mainz.mpg.de/smocc/
Notes: S M O C C
Objectives
The overall goal of SMOCC is to investigate the connection between the composition and abundance of biomass burning aerosol, the lowering of the size of cloud droplets formed on this aerosol, and the climatic consequences of the resulting perturbation of cloud physics. To achieve this goal, we will:
(1) Make field measurements of aerosol and supporting parameters in the Amazon Basin;
(2) chemically characterise the aerosols produced by biomass burning, with particular attention to the organic fraction;
(3) determine the link between the aerosol chemical/physical properties and the aerosol hygroscopic and cloud-nucleating properties;
(4) model the effect of biomass burning aerosol on cloud microphysics at the cloud and regional level; (5) investigate the effect of smoke aerosols on climate dynamics and the resulting large-scale climate effects;
(6) use satellite data to detect, validate and quantify the effect of smoke aerosol on cloud properties and precipitation processes.

## Slide 31: NSF MRI Project, Arnott PI 2002 – 2006 and Navy SBIR
Develop Multiple Wavelength Photoacoustic Instrument and Add Scattering Measurement
Navy Issues SBIR For Multiple Wavelength Photoacoustic Instrument Arnott Discusses SBIR with Sunset Labs, Rupprecht and Patashnik, Aerodyne, Spec Inc, and DMT
DMT Agrees to Commercialize Single Wavelength Instrument and Work on SBIR for 3 Wavelength Instrument Photoacoustic Instruments Complement the DMT SP2 Instrument

## Slide 32
Aerosol Absorption and
Scattering Coefficients
355 nm
405 nm,870 nm
532 nm
671 nm, 676 nm
1047 nm
Aerosol Optics Instrument
f0
f0 + df
Notes: 32

## Slide 33: Images of the Dual Wavelength Instrument

## Slide 34: Fire Science Laboratory: Biomass Burning Study
Instrumentation sampling wood smoke during FLAME Chamber burns
Photoacoustic instruments

## Slide 35
Chart: 
| 0.866710891413841 | 0.869793966707346
Ponderosa Pine, mixed fuels | 1.7116 | 1.6513
Chamise | 1.1875 | 1.3237 | 1.1974
Rice Straw | 2.8091
Ponderosa Pine Duff | 2.7929
Alaskan Duff | 3.3763
So. Californian Manzanita | 1.1828
Sage & Rabbitbrush | 1.1608
Lodgepole Pine, mixed fuels | 1.8162
Utah Juniper, foilage/twigs | 0.98763
Puerto Rico Fern | 2.6892
Wax Myrtle, branches/foilage | 2.1628
Southern Pine, needles | 1.7205
Comparison of Smoky July With Lab Data
35

## Slide 36: Multiple Wavelength Instrument
Need to translate response of red and blue lasers to f0 = 1500 Hz.
Idealized
Response

## Slide 37: Multi Wavelength Analysis Like Single Wavelength

## Slide 38: Schematic of 3 wavelength Photoacoustic instrument with scattering sensor

## Slide 39: 3 Wavelength Instrument Patent

## Slide 40: 3 Wavelength Instrument Patent

## Slide 41: Single and 3 Wavelength Instruments made by Droplet Measurement Technology of Boulder Colorado.
PAS 1 Phased out in favor of the PAX instrument

## Slide 42: PAX is Born: Circa 2007
Developed to reduce component costs and expand capabilities of scattering measurements.
DMT and UNR have a joint NSF SBIR Phase 1 in 2009 to further develop the PAX.
Created 3 mobile lab versions, 2 for DRI and 1 for the Bay Area Air Quality Management District.

## Slide 43: Schematics of the then new SUPER MINI PAS

## Slide 44: Photograph of the then new SUPER MINI PAS
Later to become the PAX

## Slide 45: Mini PAS, now known as the PAX, Prepared for the Kirchstetter lab at LBL

## Slide 46: Licensing Agreements with DMT and UNR/DRI
PAS 1 single wavelength photoacoustic instrument
June 3, 2005 (DMT and DRI agreement)
PAX single wavelength instrument
PAS 3 triple wavelength instrument
January 7, 2011
Many field deployments and associated publications:
Mexico City, Sacramento, Vancouver, Reno, Las Vegas, Oklahoma, Washington
Monterey, Riverside, Los Angeles, San Francisco, Oakland, Nevada Gold Mines
…